## Sheet1
| 1 | 2 | 3 | 4 | 5 | 6 | 7 | 8 | 9 | 10 | 11 | 12 | 13 | 14 | 15 | 16 | 17 | 18 | 19 | 20 |  | characters |  | format | units | range
1 |  | P | M | T | e | m |  | V | R | E | C | V |  | C | C | C | / | T | T | T | Tempr Charger/Coil | 4 | 5 | XX.X | C | 10.0-50.0
2 | [20 empty] | Tempr CoilTop | 4 | 5 | XX.X | C | 10.0-50.0
3 | [20 empty] | Tempr CoilSkin | 4 | 5 | XX.X | C | 10.0-50.0
4 | [20 empty] | Coil Drive Voltage | 4 | 5 | X.XX | V | 0.00-9.00 | value can reach 20.0V based on HW, or 10.0V with current register settings. No user currently uses higher than (user max to support 95mA charge are 6.7-8.25V, most below 7.5V)
[22 empty] | 4 | 6 | XXXX | mV | 0-9000
[21 empty] | Coil Drive Current | 4 | 5 | X.XX | A | 0.00-5.00
[22 empty] | 4 | 6 | XXXX | mA | 0-5000 | value can exceed 5000, but that stops charging
[21 empty] | System Current | 4 | 5 | X.XX | A | 0.00-2.50
| 1 | 2 | 3 | 4 | 5 | 6 | 7 | 8 | 9 | 10 | 11 | 12 | 13 | 14 | 15 | 16 | 17 | 18 | 19 | 20 |  | 4 | 6 | XXXX | mA | 0-2500 | value can exceed 2500, but that stops charging
1 |  | c | o | i | l | T |  | c | o | i | l | V |  | c | o | i | l | I |  |  | VREC | 4 | 5 | XX.X | V | 0-30.0
2 | [20 empty] | commanded current | 3 | 5 | XXX | mA | 0-100
3 | [21 empty] | 2 | 4 | XX | mA | 0-99
4 | [20 empty] | target current | 3 | 5 | XXX | mA | 0-100
[22 empty] | 2 | 4 | XX | mA | 0-99
[21 empty] | measured current | 6 | 8 | -XXX.X | mA | -120.0-120.0 | tends to measure ~10% higher than commanded until constant voltage taper current
[22 empty] | 4 | 6 | -XXX | mA | doesn't need decimal
| c | O | I | L |  | ! | X | X | . | X | C |  |  |  |  | X | . | X | X | V | measured Voltage | 4 | 6 | XXXX | mV | 2500-4100
| I | M | P | L |  | ! | X | X | . | X | C |  |  |  |  | X | . | X | X | A |  | 4 | 5 | X.XX | V | 2.50-4.10
| V | R | E | C |  | ! | X | X | . | X | V |  |  | X | X | X | / | X | X | X | Tempr PM | 4 | 5 | XX.X | C | 10.0-50.0
| B | A | T | X |  | X | X | X | X | m | V |  | - | X | X | X | . | X | m | A | TemprB1 | 4 |  | XX.X | C | 0.0-55.0
[21 empty] | TemprB2 | 4 |  | XX.X | C | 0.0-55.0
|  | X | X | . | X | C |  |  | X | . | X | X | V |  |  | X | . | X | X | A | TemprB3 | 4 |  | XX.X | C | 0.0-55.0 |  |  |  | CT
| V | R | E | C |  | X | X | . | X | V | ! |  |  | X | X | X | / | X | X | X | [9 empty] | 4.7 | 999
| T | E | M | P |  | X | X | . | X | C | ! | [18 empty] | 6.5 | 688
| B | A | T | # |  | X | X | X | X | m | V |  | - | X | X | X | . | X | m | A | [9 empty] | 10 | 0
|  |  | X | X | . | X | C |  | X | . | X | X | V |  | X | . | X | X | A
[6 empty] | C | O | I | L |  | O | N | [15 empty] | 4.7 | 999 | 311 | 1.8 | 172.7778
|  |  |  | U | S | E |  | C | A | U | T | I | O | N | : | [12 empty] | 7 | 601.1
|  | N | o |  | P | M |  | M | o | n | i | t | o | r | i | n | g | [10 empty] | 7.5 | 514.6
[28 empty] | 8.25 | 384.85
| e | x | t | c | o | i | l |  | ! | X | X | . | X | C |  | X | . | X | X | V | [7 empty] | 9.5 | 168.6
| i | m | p | l | a | n | t |  | ! | X | X | . | X | C |  | X | . | X | X | A | [7 empty] | 10 | 82.1
| V | r | e | c |  | ! | X | X | . | X | V |  |  | X | X | X | / | X | X | X
| B | a | t | t |  | X | X | X | X | m | V |  | - | X | X | X |  | m | A
| e | x | t | c | o | i | l |  | ! | X | X | . | X | C |  | X | . | X | X | V
| i | m | p | l | a | n | t |  | ! | X | X | . | X | C |  | X | . | X | X | A
| V | r | e | c |  | ! | X | X | . | X | V |  |  | X | X | X | / | X | X | X
| B | a | t | t |  | X | X | X | X | m | V |  | - | X | X | X |  | m | A
| c | O | I | L |  | ! | X | X | . | X | C |  |  |  |  | X | . | X | X | V
[18 empty] | X | X | A | bar: n to g
|  | N | o |  | r | e | s | p | o | n | s | e
|  |  | f | r | o | m |  | i | m | p | l | a | n | t | !
| c | O | I | L |  | ! | X | X | . | X | C |  |  |  |  | X | . | X | X | V
[16 empty] | X | . | X | X | A
|  | N | o |  | r | e | s | p | o | n | s | e
|  |  | f | r | o | m |  | i | m | p | l | a | n | t | !

## Sheet1 (2)
c | o | i | l
i | m | p | l
v | r | e | c
b | a | t | t